DAx22000

&

DAx14000

Driver for LabVIEW/Windows

User’s Guide

© Copyright ACQUITEK SAS. 2016

1 bis rue Marcel Paul

Massy

91300 MASSY

Tel: + 33 (0) 1 60 13 52 73

Fax: + 33 (0) 1 60 03 68

info@acquitek.com

www.acquitek.com

MS-DOS and Microsoft Windows are trademarks of Microsoft Incorporated.

IBM, IBM PC, IBM PC/XT, IBM PC AT and PC-DOS are trademarks of International Business Machines Corporation.

LabVIEW is a trademark of National Instruments Corp.

Changes are periodically made to the information herein; these changes will be incorporated into new editions of the publication. ACQUITEK may make improvements and/or changes in the products described in this publication at any time.

Table of Contents

1	Preface	4

2	Installing the DAx LabVIEW VIs	5

3	Overview of DAx Sub VIs	6

3.1	DAx_GetNumCards	6

3.2	DAx_Open	6

3.3	DAx_Close	6

3.4	DAx_Initialize	7

3.5	DAx_SetClkRate	7

3.6	Dax_SelExtTrig	8

3.7	DAx_Run	8

3.8	DAx_Close	9

3.9	DAx_SoftTrigger	9

3.10	DAx_Place_MKR2	9

3.11	DAx_CreateSingleSegment	10

3.12	DAx_CreateSegments	11

3.13	DAx_LoadSegments	12

3.14	DAx_PWR_DWN	13

3.15	DAx_Ext10MHz	13

4	Programming Example	14

# Preface

This manual is meant to serve as an aid to engineers using the DAx22000/14000 Arbitrary Waveform Generator in the LabVIEW  for Windows environment.

Throughout this manual, it is assumed that you are familiar with the LabVIEW graphical programming environment. If you do not feel comfortable with LabVIEW, it is highly recommended that you go through the Getting Started with LabVIEW for Windows manual supplied to you by National Instruments, before starting any program development for the DAx.

It is also assumed that you have correctly installed the AWG DAx driver.

It is also assumed that you are familiar with the IBM PC and Microsoft Windows.

# Installing the DAx LabVIEW VIs

Before installing the DAx LabVIEW Driver, you must first install and configure the DAx22000/Dax14000 Windows drivers.

To install the DAx LabVIEW Driver, simply unzip the setup in any directory. It will just copy The DAx Sub Vis and demo.vi in that directory

In this working directory, you will see three DLL:

For 64-bit OS:

- Dax22000_lib_DLL64.dll = the main DLL from the API (see all details under the DAx22000-DAx14000 manual)

- Acq_DAx_64.dll.  this DLL is only called in the DAx_CreateSegment.vi and DAx_LoadSegments.vi

- Ftdxx64.dll : the FTDI USB driver

For 32-bit OS:

- Dax22000_lib_DLL32.dll = the main DLL from the API (see all details under the DAx22000-DAx14000 user manual)

- Acq_DAx_32.dll.  This DLL is only called in the DAx_CreateSegment.vi and DAx_LoadSegments.vi

- Ftdxx32.dll : the FTDI USB driver

# Overview of DAx Sub VIs

## DAx_GetNumCards

Description

Returns number of DAx systems detected. Please note that all DAx systems must be powered up for this to occur. The DAx is not powered over the USB bus.

Parameters

n/a

Return Value

Number of DAx systems detected on USB bus.

## DAx_Open

Description

Opens API driver so that other commands may be used.

Parameters

CardNum: 1 <= CardNum <= 4

Return Value

0 if successful.

## DAx_Close

Description

Closes API driver. This is usually called before user application is terminated to prevent memory leakage.

Parameters

CardNum: 1 <= CardNum <= 4

Return Value

0 if successful.

## DAx_Initialize

Description

This call initializes the USB peripheral controller and all the internal registers on the DAx card. It must be called after DAx Open.vi

Parameters

CardNum: 1 <= CardNum <= 4

Return Value

0 if successful.

## DAx_SetClkRate

Description

This call sets the sampling rate of the D/A's. The user can choose anything between 25 MHz and 2.5 GHz as a target value for DAx22000 and 4.0 GHz for DAx14000 and returns the actual value it was able to set. The DAx22000 uses a Fractional-N Synthesizer so the results are

usually within 1ppm.

Parameters

CardNum: 1 <= CardNum <= 4

25e6 <= Freq <= 2.5e9 for DAx22000 (target value)

25e6 <= Freq <= 4.0e9 for DAx14000 (target value)

Return Value

Actual value the synthesizer was able to set ... usually within 1ppm

## Dax_SelExtTrig

Description

When the ExtTrig value is set to "true", the DAx card uses the external trigger input for all triggered modes except for soft triggers or starting trigger which the user can set at any time.

Parameters

CardNum: 1 <= CardNum <= 4

ExtTrig: true = external trigger enabled; false = external input ignored

Return Value

0 if successful

## DAx_Run

Description

This call is performed after all other setup calls have been done. It tells the card to start outputting waveforms. If TriggerNow is true, then the card creates a soft trigger once to get things going.

Parameters

CardNum: 1 <= CardNum <= 4

TriggerNow: true => add software trigger to start things off. Otherwise the trigger must be supplied somewhere else like software trigger or external trigger.

Return Value
0 if successful.

## DAx_Close

Description

This is the opposite of DAx_Run. It basically shuts off all output signals.

Parameters

CardNum: 1 <= CardNum <= 4

Return Value
0 if successful.

## DAx_SoftTrigger

Description

This is a software trigger and works whether or not the external trigger is enabled.

Parameters

CardNum: 1 <= CardNum <= 4

Return Value

None

## DAx_Place_MKR2

Description

This function call allows the user to pick one point during the waveform play time to output a positive pulse on TTL output Marker #2. The user must calculate the value by picking a point in the waveform (a time) and dividing that by 16.

16 is the resolution of the comparator. The width of the pulse is 1/16 of the sample rate (e.g. 8ns at 2.0 GS/sec).

Parameters

CardNum: 1 <= CardNum <= 4

Mod16_CNT: 0 <= Mod16_CNT <= (max memory)/16

Return Value

None

## DAx_CreateSingleSegment

Description

This call provides the simplest way to create a continuous waveform looping on itself

(NumLoops = 0), or a single triggered waveform.

The user must be sure to pass the size of the waveform (NumPoints), the number of times to repeat the waveform (NumLoops), a pointer variable pointing to the user array containing the data (UserArrayPtr), and finally, whether the segment will be self-triggered or triggered by an external signal (Triggered).

Parameters

CardNum: 1 <= x <= 4

ChanNum: 1 <= x <= 2

NumPoints:

48 <= x <= (Max Memory - 1024) [must be modulo 16] for DAx22000

128 <= x <= (Max Memory - 1024) [must be modulo 32] for DAx14000

NumLoops: 1 <= x <= 65,534 (0 = continuous loop)

PAD_Beg: 0 <= x <= 4095

PAD_End: 0 <= x <= 4095

Buffer: Pointer to an Array of type WORD (unsigned short, U16).

TrigEn: 0 => Waveform starts up first time but cannot be triggered later

without shutting down first.

1 => Allows waveform to be triggered more than once while running.

Return Value

None.

Raw data conversion: 
     0 	 min output voltage

2047 	 zero  (0V)

4095 	 max output voltage

## DAx_CreateSegments

Description

This function allows the programmer to create multiple segments and define the characteristics of each segment.

User can create up to 60 segments. After every segment has been created, he can call the DAx_LoadSegments VI to upload these segments on the generator. See this VI after.

Both DAx_CreateSegments and DAx_LoadSegments sub-VI are a mixed of the DAx22000_CreateSegments C function from the API (DLL). Due to the complexity of some passed parameters in this function, a new wrapper DLL (Acq_DAx_64.dll) has been created. Please review also the DAx22000-DAx14000 user manual for DAx22000_CreateSegments description.

For each segment, user needs to call this VI

See Multiple_Demo.VI for a better explanation.

Parameters

SegNum: 1 <= x <= 60

SegmentPtr: Pointer to User Data (Type = unsigned short, U16)

NumPoints:

48 <= x <= (Max Memory - 1024) [must be modulo 16] for DAx22000

128 <= x <= (Max Memory - 1024) [must be modulo 32] for DAx14000

NumLoops:  Number of extra times to repeat current segment

TrigEn;  If > 0 then wait for trigger before next segment.

Loop: 1 allows it to loop entire string of segments.

## DAx_LoadSegments

Description

This function allows the programmer to create multiple segments and define the characteristics of each segment.

User can create up to 60 segments. After every segment has been created, he can call the DAx_LoadSegments VI to upload these segments on the generator. See this VI after.

Both DAx_CreateSegments and DAx_LoadSegments sub-VI are a mixed of the DAx22000_CreateSegments C function from the API (DLL). Due to the complexity of some passed parameters in this function, a new wrapper DLL (Acq_DAx_64.dll) has been created. Please review also the DAx22000-DAx14000 user manual for DAx22000_CreateSegments description.

See Multiple_Demo.VI for a better explanation.

Parameters

CardNum: 1 <= x <= 4

ChanNum: 1 <= x <= 2

NumSegments: 1 <= x <= 60

PAD_Beg: 0 <= x <= 4095

PAD_End: 0 <= x <= 4095

Loop: 1 allows it to loop entire string of segments.

Return Value:

0 if successful.

## DAx_PWR_DWN

Description

This call powers down the DAx instrument

"DAx_Initialize" will start it up again.

Please note that after initialization the user must reload all settings since they are lost during power down.

Parameters

CardNum: 1 <= x <= 4

Return Value

0 if successful.

## DAx_Ext10MHz

Description

This function call must be called AFTER an optional call to "DAx22000_Initialize". When (Enable = 1) the DAx22000

expects a 10 MHz external clock reference signal. Maximum amplitude = 3Vpp; min = 0.7Vpp. For low jitter the clock

should be squarewave not sinewave. When (Enable = 0) then the DAx22000 uses the internal 50 MHz clock.

Parameters

CardNum: 1 <= x <= 4

Ext 10Mhz: Enabled for anything greater than 0.

Return Value

0 if successful.

# Programming Example

Demo.VI

Multiple_Demo.VI